## Sheet1
PRIMARY LABEL | SECONDARY LABEL | TERTIARY LABEL | FUJI XAIR |  |  |  |  | DELFT LIGHT | DELFT ULTRA | HDT MINE2
|  |  | Lunit (offline) | Lunit (box) |  | CAD4TB (online) |  | CAD4TB (box) | CAD4TB (box) | Lunit (offline)
|  |  | Pakistan | Vietnam |  | Zambia |  | Nigeria | Ghana | Cambodia
|  |  | Director | radiologist or clinician | technical advisor, m&e director | Project lead | Radiologist, clinicians | Coordinator, Director, Manager (hands-on) | Director, Radiographer | PI, project field supervisor, researcher
X-ray
| Hardware - intervention char: adapt | supporting frame |  |  |  | was given a robo-arm generator stand. Detector and generator stands described as being difficult to move 'the radiologist finds screening both children and adults inconvenient because it's a very manual system'. tradeoff between portablity and ease of operation. However the tradeoff compared to using an automatic stand like in the hosiptail setting, is the portablity. Light weight stands are usually possible only in manual format. Operating manual stands can be time consuming when screening both adults and children. rigidity of the stand might be needed, -very ill patient might cause breakage. | The manual operation of the tripod stand is awkward e.g. altering height for taller patients and lowering for kids. | 2 computers and CAD4TB box.
| complexity | displaying device | carry 2 laptops- require separate for lunit. Lunit laptop cannot be used for other software. The lunit laopto now is been replaced by X-box |  |  | 2 laptops + 1 offline CAD box. | patient demographic info and registration occurs on fuji laptop - the CAD score is viewed on the dedicated CAD laptop. |  | Has a built-in workstation in the camera which can be used to display the CXR and CAD results. however the screen on the Xray camera is small. It can also be connected to additional device through remote control asoftware e.g. TeamViewer. CAD4TB results are not shown on ultra. Compatible with laptops and tablets. The tablet screen resolution is higher. | laptop
| Set-up and connection: packaging | detector to console and to AI device |  | had problems with the bluetooth connection between the detector and laptop - occassionally connection is lost or would take a long time to load the image. the wire connector provided by the company was also too short (1-2m) for use in the field. |  | it stores them locally and what it does is when you connect to the internet it uploads them. Sometimes we find missing CAD scores, because the internet cut out for a few hours on a day they had internet problems. | Poor network connection or that the wireless transmitter from the console can lose connection can resultin failed imaage tranfer between the console and alptop. | The transfer of images from the console to the CAD laptop is sometimes interrupted until the re-start CAD can resolve the problem | possiblity of wireless connection to the detector for all the products (detector) | after taking CXR push the 'send' button from the MINE 2. no difficulty with transfer.
| packaging | assemly time | Takes around 5 minutes. |  | 10-15 minutes. |  | Assembly took 10-15 minutes. | Daily set up takes 30-35 minutes. Also clean up, taking around 35-40 minutes, using paper towels and wipes to clean equipment because locations aren't hygienic.
| varying body size | adaptivity | offer a wide range of exposure setting for patinet with varying size and weight (incl children) |  | offer a wide range of exposure setting for patinet with varying size and weight (incl children) | offer a wide range of exposure setting for patinet with varying size and weight (incl children) | careful positioning was used to take images from small children and from ill patients but field team was successful in doing this. |  | unable to take images from a very heavy pt (over 100kg) because of the fixed exposure settings of Ultra. | used different power settings. Affected image quality.
| radiation safety and regulation_outer setting, external policy |  | PNRA (Pakistan Nuclear Regulation Authority) usually don't give out licenses unless the room has lead walls, but this isn't possible in the field. They reviewed the category and tested out the radiation that was emitted using dosimeters according to the recommended saftey setting, x-ray is at least 2m away from the technician, patient is at least 2m away from generator. If closer, technician wears a lead apron. PNRA confirmed that it can be operated without a lead wall. Require radiologist to wear dosimeter to check exposure after 1 year. the manufacture is resposniblt to obtain a certificate to import the device, then implmeeitng partners have to get a certificate to operationalize the device. |  | device had to be registered for import (this was done by fuji), studies were also done in the National Lung Hospital checking how the image quality and radiation exposure compared to other xray systems. System is now registered for national use. The device has to be set up far away from the crowd. Operator operates xair at a distance and detector is 3-4 m away from xair. | similar to pakistan. Notified the NTP. Sought permission from the radiation protection authority of Zambia, who made an assessment re its use in rooms without lead lining. Initial permission for import was based on the equipment dosier. After arrival, equipment was assessed on site for radiation leakage. device passed and was allowed to use in non-specialist room. got license to use it. Radiographer required to wear radiation badge. Radiographer wears radiation apron if operating in room with patient. Or uses a Xair cable to take exposures outside of the room. Set-up is in an isolated room in facility with closed doors. | Warning signs for patients (esp pregnancy onces). Feel safer because Xair is low dose. Lead apron when unable to leave distance between xray operator and generator. might use lead wall and glass. | delft responsible for import regulation. Had already secured approvals for a previous project from the nuclear regulation agency and this was extended to include the Delft Light. wear a lead radiation jacket and collar on. Maintain a distance from the generator using the handswitch, ensure no one is between the generator and detector. 'site selection is important because we can't operate inside a lead wall'. Barricade certain areas to make sure patients don't enter radiation zone. no one goes behing the detector for a distance of 10m. | just began process of device registration. Lead gown and shield, stood 2m in a diagnoal fashion away before making the exposure. | radiation shield, two lead aprons, dosimeter included with HDT.
| Power and electricity- intervention char: adapbility |  | can't change and takeing exposure at the same time (foot pedal to switch between charging and exposure modes). If battery is insufficnet, things that require mains electricity: generator, DR battery, fuji laptop, lunit laptop. average 35 shots on battery. Limitation of throughput (see challenge) |  | average X shots on battery. can't change and takeing exposure at the same time (hand switch is used to between charging and exposure modes). This affected throughput handswitch for the geenerator is provided. battery life was improved with the capacitor, but they still opted to use a second x-ray that was stationary to compensate for this reduced throughput. | not meant for high throughput on continuous use. |  | detector battery is switched in order to operate for one whole day. But the generator battery cannot be swithched but is durable and can be used for a whole day continous operation. | Ultra - battery powered - cannot be use to take exposure and charge at the same time. CAD4TB box powered from the router. | electrical mains or battery
|  | external | Spare detector battery provided, but swapping batteries requires disassembling the panel which takes time (around 5 mins) affecting throughput. |  | capacitor' (power bank) was purchased locally by fuji in response to battery problems. By using the device in this way, you can use the device over a whole day or multiple days. | solar panel and power pack for generator provided in response to battery issues. These extended battery life so they could screen for the whole day. Issues with solar panel recharging using sun light. See challenges |  | power bank and solar panel. See challenge / malfunction | Batteries can be swapped for the detector but not generator. | generator by gasoline, but not had to use so far in field.
CAD software
| Internet requirement_adaptbility |  | reading is done offline through a lunit laptop |  | CAD reading done offline in the X-mobile box | reading done using the offline box,a nd the upload to the cloud using online (hydrid) using Delft cloud system. Images are stored locally when operating without connection and then upload to cloud occurs when internet connection is restored, delaying reading. Occassionally delayed upload is possible. |  | Both [online and offline]. Same as zambia, hybrid. The system has a router attached to it and in some locations there is no internet, but they can continue with the screening and upload to the cloud occurs when they're able to connect to the internet. | only use offline CAD4TB box. | offline with laptop. using CXR computer as server
| interpretation/use case --> describe the dataset |  | using the score for screening and traige |  | research | research. Concern of using CAD as the ground truth. | research | screening and traige | triage | screening and triage (using heatmap output and score).
| Threshold score selection_ Process executing |  | remained at a default threshold score (50). Considering changing threshold once more results have been collected. |  | felt to depend on goal. 'We operate in an environment that's not very Xpert cartride rich, so we try to negate additional testing but have good specificity as well'. Noted that this depends on the AI product you're using. E.g. Delft recommend a default, Qure.ai have a system of autoselection using project results. | did not select a threshold score for this implementation. In the past, made mistake where used the same threshold for two projects then data analysis showed the optimum threshold was substantially lower (45 versus 60) but this was not too critical because using it in parallel with symptoms. in process of reviewing CAD guidelines with NTP and 'I insisted that we need to put in that the CAD threshold has to be properly selected and then given to the teams to use. It is not something to just choose.' Has concerns about general use of CAD and threshold selection in case clinicians interpret CAD score as indicating TB absolutely, and suspects this already occurs in Zambia [see quote in document]. |  | Use a CAD4TB score of greater than 60 to triage. This was manufacturer-recommended, and they had used successfully in previous projects. Initially, a score of 40 was used but yield was low so they increased. 'I think the structural defects are very important in score selection'. Do not adjust the score in the field - rely on clinician clinical judgement to capture patients with suspected TB but low score. 'We don't stick strictly do 60+, we just use this as a quick judegement when doing population screening. Symptomatic patients with a score of less than 60 are still screened'. | threshold score was already decided by Ghana NTP. 60. this threshold was decided based on prevalence screens. | used recommended score. estimated this to be appropriate in terms of tradeoff of sensitivity and Xpert cartridge and sputum collection resources.
Package | x-ray and CAD_Intervention char: complexity | displaying device | pre-isntalled and integrated | pre-isntalled and integrated | pre-isntalled and integrated | didn't use the fuji-lunit package but insteaded used FUJI -CAD4TB. therefore needed another laptop to link the Fuji and CAD4TB systems, to provide the link between the cloud and CAD4TB. This was necessitated because the fuji laptop was unable to install other programs like Google Chrome. Integration was easy and was collaboration between mannufacturer and project IT teams. | Marrying the Xair and the CAD4TB, they're working well and we're able to get the results we want.' | pre-isntalled and integrated | pre-isntalled and integrated | pre-isntalled and integrated
| storage/ data sychronization-executing |  | lunit laptop storage space was a problem - prevented report generation and printing. All data is backed up on 2 hard drives, after a camp is compelted the DICOMs are saved from the Xair laptop. When getting closer to storage capacity, lunit is slow. |  | Dicom back ups were taken from laptop onto a hard drive for storage and dicoms deleted off the laptop. This is necesitated because laptop storage fills up and it starts deleting data. | CXR are uploaded to Delft cloud for analysis and for storage. Cloud was used as a consulting platform - to allow non-field-based staff to access CXR for second-opinion. This was key for project site. |  | Can extract high resolution DICOMS. Transfer to server only occurs online. | CAD4TB box itself stores images. Exposure can also be stored on Ultra camera. During the screening pen drives weren't used but this is possible.
| Other accessory |  | maybe mentionign printing on paper. |  | results were reported to patients and written down. Abnormal films were printed. |  |  | Use a barcode machine to register patients .
| Privacy - executing |  |  |  | wrote programs to extract metdata from patient images, use a system with unique patient identifiers. |  |  | Anonymization before upload. | Mentioned only about extracting and storing DICOM, the need to follow data protection regulations of Ghana (and the necessity of this for any country). | secure server. De-identification of data using a patient ID system. KNTA local server stores the file.
Implementation
| HR_Inner setting: Readiness for Implementation, Available Resources |  | 1 district coordinator per district + 2/3 field officers are the project employed staff for screening camps. Government employees support. For mobilizations CHWs (2/3 per 200 people) are employed from the local area. 2-3 trained x-ray technicians on board, these are fuji trained governemnt employees. 1 per screening camp. Core: coordinator, trained xray technician(s), chws (directing pts). used CAD to replace radiologist therefore no clinician/radiologist onsite. |  | For tb, usually have a couple of screeners (3-5) depending on the volume of the event. 2-3 people operationalize xair. Also had additional staff who screened for non-TB diseases (diabetes, etc). | core operational team (similar to others) plus research staff centrally-located to provide support. Radiographer receives CAD output, clinician would make the triage decision but the radiographer has also undergone training and can leave comments on a CXR with reference to the CAD score and CXR image. |  | 3 people on project field team: 1. Team coordinator, 2. radiographer, 3. data clerk (dedicated to registering pts). Medical doctors and radiologists read the images. Radiographer operates the CXR. Needs at least two people to operate the equipment. | Core operational staff (+ extra clinicians and IT officers). Total 5 staff. Some staff were from NTP. IT specialist was needed during installation and integration (IP addresses). | core project team, + additional staff: data manager, driver, receptionists from community. IT team responsible for installation.
| Training_Inner setting: Readiness for Implementation, Available Resources |  | Fuji provided training and this education was generic (what fuji has done over the years, what the device is and how it works etc). Only a general explanation of how CAD works, but no background. 2 days, one day focused on calibrating and operating the device. Hands-on section of training was more valuable. Training described as good and no issues post training. Additional hands-on learning from the Singapore office for how to fix small bugs would be helpful. |  | Thorough and supportive. On-site training soon after equipment delivery. Training described as thorough and supportive. | initial training was remote, online from s.africa fuji office. This happened soon after equipment arrived. Interactive. Explaining equipment set-up and recommended parameters. Training was 2 days or whole day. No further training needed or difficulty using machine. | no problem experienced after training | training was online, project staff had difficulty using the device after training so follow-on call-in trainings occurred to assist field work. Think physical training would be more appropriate. | [manufacturer] | by manufacturer and internally. Internal training was via online lectures. Also did practical training. No difficulty after.
Manufacturer Input |  |  | Singapore team is well trained, local Pakistan team is not well trained and rarely solves queries independently (singapore team provide remote technical support). 'The guys in the Pakistan office, especially the technicians, are not really well-trained. If there is a mass implementation of Xair, say the country gets 100 units, they will have to employ more qualified and proficient people,' . |  | fujifilm have a local office in Vietnam. Very positive about experiences with fuji team. | though fuji have local office in south africa - when equipment crashed it had to be sent back to japan. [would like to know how well trained the south africa team is]. Fuji has already entered the Zambian market and have a local distributor. |  | assisted in the anonymization of data with the project team.
| Maintenance_Readiness for Implementation, Available Resources |  | Fuji supported installation and integration of the CAD and ultra-lightweight CXR. As a pilot project, DOPASI received lots of support from Fuji. Generator needed replacement and this caused significant delay as coming from Singapore (delayed screening). |  | too early in implementation Because of covid-19 a lot of IT support has been remote, and this might be part of the reason why they've found set up so difficult. Fuji has visited project sites in person but a lot of support can be done via the internet. | Like Pakistan, maintenance and troubleshooting is prompt and supportive. CAD and DXR integration was a collaboration of IT specialists at CIDRZ, Delft and Fuji. |  | Installation was guided remotely. 'It took a longer time for us to be able to put together the Delft Light package, despite our prior experience and experienced radiographer staff members. these things could easily be done in one day, with physical training, but online training prolonged this experience and it took longer than a week to figure this out'. Engineer is available online 24/7 'if you go to the field without internet connection, you have to wait until network is available to be able to talk to the engineer. All of this could've been avoided using physical training'. | [1 day] | real-time technical support and part exchange.
| Importation -external policy |  | Securing permission to import. Manufacturer borne the costs associated with import: customs, freight. |  | local office took care of device registration and import. Loaned them second machine when waiting for domestic approval for new machine. | check transcript |  | Responsible for regulatory requirements for import. |  | complex. Issues experienced. Took at least 4 months.
Cost |  |  | both CAD and Xair were free. Cheap printing option chosen because did not budget for CXR printing. 'Obviously, it's a cost effective solution'. |  | I think most people are familiar with the pricing models of a bundle price or a price per read, or a price per bundle of reads, but how does that integrate into the program when you have a radiologist who is not going to be replaced by the software.' Think the device is expensive at a cost-per-case detected. | USAID were interested in performance because DXR is cheaper than other x-ray systems and they could consider procuring more of these for Zambia. |  | it is less expensive as you do not need any chemicals, everything is digital'. | unsure of cost but cited as being cheaper than the other machines they have.
User Experience |  |  | the portability, the outreach, the access to services for vulnerable population segments is the best thing I can think of'. | Good quality and easy to set-up, but battery throughput is an issue for our campaigns'. | xair was noted as being much more convenient for their target population, compared to a truck the throughput is faster and more streamlined. Allows you to utilize outdoor space. 'From my perspective, the team seems very excited to have that in the field and we're happy with how it's used'. 'As a downside, I think there's still a bit of hesitancy about bringing in foreign products, about monopolies and pricing, and about the role of the radiologist in the CAD program'. 'The Lunit INSIGHT box has come with more issues than we initially anticipated. It's been more of a challenge to set-up, but the team has been incredibly supportive and helpful'. | It’s so portable that we can move it around and the image quality is good; but, for high throughput screening it poses some challenges'. 'It's really just the battery thing. I understand that it's a very small unit, so the battery is small and there's only so much you can ask. I guess in the future, as the technology changes, and other batteries become available, then this will become perfect'. Recognized the utility of the device for screening rural communities. Could wheel the system in the case they were provided with or could lift it with one person. Also felt packaging protected device during transport. | Coupled with experience, and the recent trainings we've had, it's a handy tool but you have to keep in mind the pitfalls of using it'. 'I think it's the way to go. It's very efficient, easy to handle, and you can take it anywhere'. | We are very impressed with the results we are getting'. 'People in these rural locations don't have access to medicine, and use traditional medicine instead, so we have to transport everything in the delft light backpack to grant those people access to care with the screening tool'. [CAD] 'this technology is helping in screening target populations, even the general population, for TB and even other diseases'. [radiologists] 'a miniature machine doing what giants cannot do'. | It would be good if they allow for more [power] options, but as it is it's very good.'. Would also like to be able to switch generator batteries to extend battery life. 'Very good for triaging TB cases, it helps accelerate the number of cases we find. For example, in our setting, the analysis is ready after less than 1 minute (around 40 seconds) and this helps us screen more people than having only radiologists'. Ultra waiting time cited as the only drawback by the radiologist.
| Image quality_4_Individual char_Knowledge & Beliefs about the Intervention / 5 process evaluating & reflecting |  | Based on my own experience, the image quality, despite the low dosage is extraordianry.'. Xair laptop allows image quality to be enhanced with software and this is quoted as a benefit. Radiological professionals from tertiary hospitals in Islamabad also impressed. thought to be same quality as stationary machines. | good | before deployment, image quality was assessed at the national lung hospital and they ceritifed that it was good. | image quality is good. 'fantastic'. | a major advantage of the Fuji software package was felt to be the ability to manipulate the CXR image to improve quality. Able to get good images from children. Image quality was felt by all three interviewees to be comparable to or better than previous analogue and stationary systems. | no difference noted with normal x-ray. 'the image qualities are very good, you cannot differentiate from that of bigger machines'. [patients] 'those that have done x-rays in the past are simply surprised that this can be done at their doorsteps with a small machine'. | images presented on generator screen can be rather small (but depends on user preference). 'The delft ultra images were good enough for the setting and good enough for TB control. I think I recall there were some motion artefacts from the heart, but not the patient, only the heartbeat. For TB screening, this wasn't a problem at all.' Comparable to Light. | suggest to be lower for larger patients. 'Does not fit for persons being around over 80kg'. Diaphragm blurring when elderly pts breathe.
| Portability_intervention char: adaptability |  | Portability depends on the weight/ dimensions of all the system. lunit laptop is old and heavy. The packaging for transport is inefficient and means its not able to be transported by one person. The project team use a heavy steel detector stand that is not easily mobile. Currently requires 3 people to transport. 'I have given demonstrations of this device to the TB NIH and the leaders of the NTP; everyone is quite amazed by the size and portability'. The equipment isn't too heavy to be carried by one person, but the packaging is the problem'. | one staff can carry the equipment but it is easier to split it between two. | We have previously worked there and would usually take an x-ray tuck, with a stationary x-ray machine, to the island. But, because of the time of year, this was not possible, so we brought over the ultra-portable x-ray machine'. Noted that portability depends on set-up but theoretically one person could carry everything (they used two people to carry it). | the project team was requested to take the device to another district with high tb burden (~5km away) to provide x-ray services there and they were able to do that because of the portability of the unit. | unit comes in a robust package that fits the computer, generator, and detector. The stands are external. The project team felt one person could manage transportation, but two would be better. Easily moved it between the two project sites. | not portable enough to be carried by a single person. 'The DR panel, the generator, and the battery are the only things contained in the backpack. The solar panel is in another bag. You also need a detector stand when you get to the field - this is large and requires its own bag'. No longer use the solar panel because the powerbank is the more portable option. Total weight 70kg. There are lots of components in the field - it's not that simple. Might use a motorcycle to take it some places. However, still more portable than the truck (previous project). 'the portability of this equipment makes it easy to reach people that needs it'. | the ultra is more portable than the light, but they both produce images of good quality'. 'of course it is, very portable, enough for a single person to carry'. Detector fits in small slim backpack with chargers and tablet, briefacse for the ultra carries the stands. ' It will be able to go to remote areas that a car couldn't: you could cross a river, or take it on a bike. It will be able to go to remote areas.' | not portable enough for one person to carry because of DXR stands and detector. Used vans for transport.
| Equipment malfunction_availbale resource |  | handswitch errors after 2 months - had to switch it on/off in between exposures. Later, another error was experienced with the handswitch. All faults were fixed or replaced by fuji. Calibration problem with lunit laptop when storage is full. The Xair generator stopped working altogether (potentially battery problem) around March - this delayed screening for 10 days while a replacement was secured. 'The screening has stopped a few times now because of minor problems like that. But, because the people here are not well trained, we face these issues and they send replacements without debate'. When heated the Xair slows down. |  |  | Fuji laptop crashed trying to integrate the systems because unable to install other programs (Google Chrome) on there. All equipment had to be returned to Japan for recalibration as a result. The detector failed once after a patient knocked it over, it couldn't reconnect with the laptop (this was resolved the same day) |  | The first solar panel supplied did not charge in the field, so this was replaced with a power bank. Battery developed a fault early on (01111) and had to be replaced - this did not interrupt field operations. | only used equipment for 1 day | equipment heats up after use for long time.
| Battery life_ intervention char: adapt |  | In general, battery-powered devices are limited by the battery lifespan and the performance/ capacity degrades with time. in general, fuji lasts less than the 100 shots claimed by the manufacturer. Particularly, the CAD laptop battery can drain faster than the console laptop and may not last the entire day of operation without charging. A solution to this is a battery powerbank which can be carried for use in remote setting (with consequences on portability of system). 36 images maximum on recommended settings. Laptop batteries are limiting in settings without power. 'The thing I'm hating the most is the battery issue. I would like it to take at least 100 shots.' Would like some external batteries for the Xair generator that could be used as a backup. design of the DR panel could be better so that battery is easier to change. | with charging, they could do around 200 exposures a day in 8-9 hours. | the operational set up is quite annoying because Xair's battery life is so short'. Achieved around 44 exposures from 1 charged battery. 'It's very useful to certain use cases but if you mobilize a lot of people in one place then the battery becomes a factor'. 'We really operate a model thats based on scale, so we really need to achieve a high throughput in terms of x-ray screens to make the investment worthwhile. Even though Xair can go to remote and hard-to-reach sites, we still need that scale coverage'. | initial unit had limited battery life, between 20-30 exposures, now on max 35 exposures before generator runs out. 'Sometimes there's no power in the facility and it runs out at 30 so we can't continue screening. That was a problem'. 'You know it was a high throughput study, we were screening up to 200 in a day, I don't think we can use the Xair for that. Having said that, I still like it because it is easy to move around' | The battery is a major weakness for the Xair'. Were only able to see 30 patients before generator required recharging. 'The battery is a problem because sometimes we have to stop seeing patients to let it recharge, then resume after an hour or so'. | Initially, battery life did not last but is improved with the new power bank. 90 is the average daily screen, using the power bank. 7 hour battery life. The generator battery is the most durable in the whole system 'it has never disappointed in the field'. The CAD laptop battery drains quickly. Clinicians like the fact that it is battery operated. | 25 on a single day. Tablet battery drains fast so use laptop instead. Without external power source, operates for 2 hours on battery power and cannot be charged during operation. | limiting in field, lasts 5-6 hours. Not enough for day
bought separately- connected with KNTA it staff. This required a lot of effort because not built-in.
Project lead | CAD | Likes | Improved capacity / Fast results / Multi-disease tool/ image quality |  | AI accuracy | Cloud platform/ image quality/ |  | AI accuracy [inter/ intra-reader variability] / Multi-disease tool/ Fast results |  | improved radiological capacity. Inter- intra-reader varaibiltiy. Fast results. Ease of use.
|  | Dislikes |  |  | Pricing structures / Foreign product/ Role of radiologist/ Difficult set up | Threshold selection concern/ |  | AI accuracy [false positives and negatives] |  | none
| DXR | Likes | Portability |  | Portability/ flow of patients [compared to truck] | Portability/ image quality |  | Portability/ image quality/ low cost |  | portability. Reduced radiation. Ease of use. Image quality [low body weight]
|  | Dislikes | Limited battery / Portability [packaging]/ Detector design |  | Limited battery/ waiting time | Limited battery |  | Portability |  | image quality [large pts].
Clinician | CAD | Likes | AI accuracy |  |  |  | Fast results / Cloud platform |  | Role of radiologist/ Fast results
|  | Dislikes | AI accuracy [false positive or negative]
| DXR | Likes | Portability/ Image quality |  |  |  | Portability/ Ease of use/ image quality | Battery-operated
|  | Dislikes |  |  |  |  | heatmap doesn’t appear on all reads
Radiographer / Radiologist | CAD | Likes | High-tech/ Fast results | Screening tool |  |  |  | High-tech | Fast results/ Multi-disease/ improved capacity
|  | Dislikes | false negatives on CAD | [6 empty] | laptop too heavy
| DXR | Likes | High-tech/ portability/ image quality/ safety/ cost effectiveness | Image quality/ ease of set-up |  |  | Ease of use/ image quality |  | Cost/ high -tech/ portability/ ease of use | portability. Image quality [most pts].
|  | Dislikes |  | Limited battery |  |  |  |  | waiting time | battery capacity. Image quality [large pts]. Slow with remote control.
| MISC |  | clinicians see utility of DXR in emergency room. Radiologist likes image manipulation software | CAD performance may differ between CXR brands |  |  | image manipulation software liked by both radiologist and clinician | issue with barcode printer |  | CXR easy to install with other devices. Speed gets slower when using remote control. Would prefer that lunit provides score separately, rather than in image tag. Pts perceive as trustworthy.
few radiology technicians around to use the machine.
Note on H6: ====== ID#AAAAW89IyDo tc={1AB9BF60-42AB-4F00-AA6C-63DA022EA2AA} (2022-03-21 08:19:11) [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: I belive someone commented on the detector mouting system. how that doens't allow to be lowered for children.
Note on H8: ====== ID#AAAAW89IyDw tc={4DDF28B9-0A1E-4A00-BCB4-ECED750AE0FE} (2022-03-21 08:19:11) [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: i recently found this note when reviewing the transcript - i think it is a different issue to the next column along
Note on K10: ====== ID#AAAAW89IyDg tc={2E5FEE15-DC41-4A00-9210-91419A1E0F79} (2022-03-21 08:19:11) [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: let's add another point - artifect is not importat for TB.
Note on L13: ====== ID#AAAAW89IyDs tc={33489330-DC64-4373-8DF7-903AF667763F} (2022-03-21 08:19:11) [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: @Rachael Barrett what does it mean by "gasoline"? Reply: petrol - i'm guessing a petrol powered generator for use to generate electricitiy, not just store it like the power banks/ solar panels